## Multiple_choice
Question ID | Source Ticket | Topic | Difficulty | Question Type | Question | Choice A | Choice B | Choice C | Choice D | Correct Answer | Technical Reality | Physical Principle | Real-World Example / Analogy | Suggested Customer Response | Key Takeaway | Status | Reviewer Comments
Q001 | TCK80900 | Metrology and scale readability | Intermediate | MCQ | A customer requests more graduations on a pressure gauge because they want a more accurate measurement. Which statement is the most technically correct? | More graduations always improve system accuracy | More graduations improve the mechanical accuracy of the gauge | Graduation design contributes to readability, but the performance of the complete measurement system must be considered | The number of graduations has no influence at all on measurement quality | C | Graduations influence readable resolution, but they do not change the intrinsic performance of the sensing mechanism. Too few divisions can limit useful reading, while too many can make the dial crowded, slow to interpret, and difficult to read consistently. The final result depends on the mechanism, scale, pointer, viewing conditions, installation, environment, and operator. | Measurement-chain uncertainty, resolution, readability, parallax, repeatability, and human interpretation. | Consider a precise load cell accurate to 0.1% of full scale connected to a display showing only integer values. The display prevents the user from fully exploiting the sensor. Now compare it with a low-cost sensor connected to a display showing four decimal places, where the last two digits constantly change. The second system looks more precise but may provide less trustworthy information. Neither the sensor nor the display alone defines the useful performance of the complete system. | Additional graduations may improve readable resolution, but only if the scale remains clear and coherent with the gauge’s actual performance. We should evaluate the complete measurement system rather than the number of markings alone. | Useful accuracy results from a coherent measurement system, not from one impressive-looking component. | Corrected
Q002 | TCK80900 | Metrology and scale readability | Beginner | True/False | Increasing the number of graduations on a pressure-gauge dial always makes the final measurement more accurate. | TRUE | FALSE |  |  | B | More graduations can sometimes help the operator estimate a value, but excessive graduations may reduce clarity and reading repeatability. They cannot improve the mechanical accuracy of the instrument by themselves. | Resolution is not identical to accuracy. Human readability is one contributor to the useful performance of an indicating system. | A ruler with hundreds of extremely close markings may theoretically display more information, but if users cannot distinguish or interpret the markings reliably, measurements can become less repeatable. | More graduations are useful only when they remain readable and are consistent with the actual instrument performance. | Additional information is valuable only when it can be interpreted reliably. | Corrected
Q003 | TCK80900 | Metrology and measurement chains | Advanced | Advanced Engineering Thinking | A highly accurate sensing element is connected to a display that only shows integer values. A less accurate sensor is connected to a display showing four decimal places. Which system is necessarily more accurate for the user? | The first system, because the sensing element is more accurate | The second system, because the display has more decimals | Neither can be selected without evaluating the complete measurement chains and their intended use | Both systems have the same accuracy | C | The user receives a result from the complete measurement chain. Sensor accuracy, signal conditioning, conversion, transmission, display resolution, stability, calibration, environment, and reading method can all introduce limitations. A precise component does not guarantee a precise indicated result, while additional decimal places do not prove accuracy. The intended use also matters because a coarse local display could coexist with a more precise transmitted signal used by a control system. | Propagation of measurement uncertainty and the weakest-link effect within a measurement chain. | A high-quality photograph displayed on a very low-resolution screen loses visible detail. A poor-quality photograph shown on a high-resolution screen gains pixels but not genuine information. The final experience depends on both the source and the display chain. | We need to define where and how the value will be used, then evaluate every element between the physical quantity and the final decision made from the displayed or transmitted result. | Always evaluate the complete measurement chain in relation to its intended use. | Corrected
Q004 | TCK84886 | Pressure range and operating envelope | Intermediate | MCQ | A process normally operates at 4 bar but can reach 6 bar during startup. What is the best engineering approach when selecting the pressure gauge? | Select a 4 bar gauge because startup is temporary | Select a 6 bar to cover both events | Select a suitable higher range (10bar) and investigate whether the peaks are slow pressure variations, pulsations, or pressure shocks | Keep the 4 bar range and add a pressure limiter below | C | The full operating envelope must be considered. A temporary condition can still overload or fatigue the sensing element. The appropriate solution depends not only on peak amplitude but also on duration and whether the event is a controlled transient or a pressure shock. | Mechanical stress, fatigue, overload, dynamic response, pulsation, and water hammer. Good practice says you should work with your operating pressure between 30 and 70% of the range. | A bridge is not designed only for average daily traffic. It must also withstand defined maximum loads and relevant dynamic events. Similarly, an instrument must be selected for credible operating conditions rather than the most common value alone. | We can first follow the good practices and offer a 10bar if few information are available. If the customer comes back with details we can confirm the choice or recommend an protection option or accessory. | Select instruments for the credible operating envelope, not only the nominal operating point. | Corrected
Q005 | TCK84886 | Pressure range and transient conditions | Beginner | True/False | A short pressure peak can be ignored when selecting a pressure gauge because the process normally operates below the gauge’s full-scale value. | TRUE | FALSE |  |  | B | Duration is only one aspect of a peak. A rapid pressure rise may create high dynamic loading, while repeated peaks may cause fatigue. The peak value, rise time, frequency, process conditions, and instrument construction must all be considered. | Transient loading, fatigue accumulation, and dynamic pressure effects. | A person may safely carry a moderate load for a long time but still be injured by a short, sudden impact. The size and speed of the event matter as well as its duration. | Even short peaks must be characterized. We need to know how high, how fast, and how often they occur before selecting the range or a protective accessory. | Temporary conditions are still operating conditions. | Corrected
Q006 | TCK84886 | Pressure range and risk assessment | Advanced | Advanced Engineering Thinking | A customer reports a 4 bar operating pressure and occasional indications near 6 bar. Which question should be investigated first before simply recommending a higher-range gauge? | Which type of protection method he would prefer to add ? | how was the peak observed ? And details about the installation ? | Is this problematic as the gauge will resist 6bar ? | does he want to replace the gauge for safety with a new one ? | B | An indicated peak can have different causes. It may represent a genuine slow pressure increase, pulsation, water hammer, pointer oscillation, vibration, temperature effects, or an installation problem. Answer A and C are not completely wrong but starting with the context you might choose the right solution instead of just going for a stronger tube | System diagnosis, dynamic measurement response, cause-and-effect analysis, and trade-offs between range and readability. | If a car warning light appears, replacing the warning light with one that tolerates a higher voltage does not identify the fault. The first task is to determine what caused the indication. | Before changing the range, we should characterize the event and confirm whether the displayed peak represents the actual process, a dynamic effect, or an installation issue. | Diagnose the phenomenon before choosing the solution. | Corrected
Q007 | TCK90540 | Diaphragm seals and fill fluid | Intermediate | MCQ | A diaphragm-seal assembly has been separated and some filling fluid has been lost. What is the correct response? | Reassemble it and adjust the pointer to zero | Add any compatible-looking oil and return it to service | Have the assembly correctly refilled, evacuated, checked, and calibrated by an appropriate facility, or replace it | Continue using it if the diaphragm is undamaged | C | The diaphragm, fill fluid, connection path, and measuring element form a closed pressure-transmission system. Fluid loss or trapped gas changes the transmission behaviour and can cause zero shift, delayed response, temperature sensitivity, or an incorrect measurement. Simply returning the pointer to zero does not restore the physical integrity or metrological performance of the assembly. | Hydraulic pressure transmission, fluid compressibility, thermal expansion, trapped gas, and calibration of an assembled system. | A hydraulic brake system cannot be repaired by simply repositioning the brake pedal after fluid has leaked. The closed hydraulic path must be restored and air must be removed before reliable force transmission is possible. | The fill fluid is an active part of the measuring chain. If the sealed system has been opened or fluid has been lost, the complete assembly requires controlled restoration and verification. | Restoring the indication is not the same as restoring the measurement system. | Corrected
Q008 | TCK90540 | Diaphragm seals and fill fluid | Beginner | True/False | If a diaphragm-seal assembly returns to zero after reassembly, that proves the system has been correctly restored. | TRUE | FALSE |  |  | B | Zero is only one point. A system can indicate zero and still contain too little fluid, trapped gas, leakage, or an incorrect fill. These defects may appear through span error, poor dynamic response, thermal effects, or non-repeatable behaviour when pressure is applied. | Calibration requires evaluation across the useful range, not only at zero. | A weighing scale may display zero when empty and still give the wrong result when a known mass is applied. Zero adjustment alone does not validate span or linearity. | A zero check is useful, but it cannot verify pressure transmission and performance across the measurement range. | One correct point does not validate the whole measuring function. | OK
Q009 | TCK81382 | Diaphragm seals and wetted materials | Intermediate | MCQ | When choosing a diaphragm seal, should all wetted parts be in the same material to resist the fluid ? | Yes, they are all in contact with the same fluid so require the same protections. | No, they are different part so they potentially require different levels of protection | the gasket should always be the point of focus to avoid external leaks. |  | B | All part in contact are in contact with the same fluid. But, they don't have the same constraints, while the flange is a solid static bloc, the diaphragm is a thin moving part, a lot quicker to damage. Similarly a gasket is constrained so it is interesting to check. | Corrosion mechanisms, galvanic interactions, chemical compatibility, permeation, temperature, concentration, and exposure time. | if you leave both a watch movement (316L) and a plate of metal (316L) outside, they'll both corrode overtime, but the watch movement will loose it's functionnality long before the plate. | the customer is in charge of knowing what is necessary forhis process. If he reuire guidance we can help but we will never have the complete process overview. | Compatibility belongs to the complete wetted system, each component is different. | Corrected
Q010 | TCK84367 | Sour service and material compatibility | Intermediate | MCQ | A customer proposes 316L stainless steel for sour-gas service because stainless steel is generally corrosion resistant. What is the most appropriate response? | Accept it because stainless steel is suitable for every corrosive service | Reject all stainless steels for every sour-gas application | Suggest Monel wetted parts which is best for corrosion. | Suggest HC276 wetted parts and gold plated diaphragm. | D | Sour service may involve damage mechanisms that are not predicted by general corrosion resistance alone. Suitability can depend on hydrogen sulfide exposure, pressure, temperature, material strength and hardness, heat treatment, stress, fabrication, and the applicable service requirements. Sour gas treatment also often imply Hydrogen induced cracking | Environment-assisted cracking, hydrogen-related damage, stress concentration, and the interaction between material properties and service conditions. | A safety rope may look clean and free from surface corrosion while having internal fibre damage that reduces its strength. A good surface appearance does not prove structural integrity. | 316L cannot be approved or rejected, sour gas can be diluted or at low pressure and temperature which is fine. But being a mortal gas for human and with the usual sour gas being high temperature, offering a good safety (gold for H2 protection + HC sour gas resistance) is important. | Corrosion resistance does not guarantee resistance to every damage mechanism. | Corrected
Q011 | TCK84367 | Sour service and material compatibility | Beginner | True/False | If a material shows no visible damages in sour service, it means the product is suitable for the application. | TRUE | FALSE |  |  | A | If something being used in real condition (for enough time) shows no damage, it is safe to say it's adapted to this environment. It's the best test. In sour gas, some damages by themselves could be invisible (H2 induced stress) but they still would overtime increase corrosion, induce measurment drift of diaphragm bubbling. | Subsurface damage, hydrogen-related mechanisms, and environment-assisted cracking. | Real condition testing is always the best test you can do. | real use and no damages is the best insurance it'll continue like that. | Real condition testing is the only way to ensure truly that all conditions are accounted for. | Corrected
Q012 | TCK89819 | Deadband and product specifications | Advanced | Advanced Engineering Thinking | A switching device has a lower deadband in controlled testing than is consistently achieved in real installations. Which specification approach is the most responsible? | Publish the best laboratory result because it makes the product more competitive | Publish a realistic value supported under defined operating conditions and explain the influencing parameters | Remove the deadband specification entirely | Allow each salesperson to choose the value depending on the customer | B | A specification should be meaningful, reproducible, and connected to defined conditions. Deadband may depend on product design, set point, adjustment direction, pressure-change rate, temperature, mounting conditions, tolerances, and test method. A more conservative but consistently achievable value gives the customer better engineering information than an exceptional result obtained under favourable conditions. | Repeatability, hysteresis, switching mechanics, test conditions, tolerances, and measurement uncertainty. | A vehicle’s lowest fuel consumption recorded under ideal test conditions is not necessarily the value a user will obtain in traffic, at low temperature, or with a heavy load. A useful specification needs stated conditions and realistic expectations. | The deadband should be discussed together with the operating conditions and test method. We should provide a value that is realistic and reproducible rather than presenting an isolated best-case result. | A trustworthy specification describes repeatable performance under stated conditions. | Corrected
Q013 | TCK91320 | Fault diagnosis and claims | Intermediate | MCQ | A recently purchased pressure gauge has been used only a few times and no longer returns to zero after depressurization. What is the best first response? |  | Continue using it and subtract the zero offset from every reading | Stop using it for decisions, document the conditions, and return it for technical assessment through the claim process | Open the case and lubricate the movement | C | A persistent zero offset after depressurization can indicate permanent deformation, movement damage, overload, shock, or another internal fault. Without controlled testing, an on-site correction could hide the symptom without restoring accuracy or safety. | Elastic deformation, zero error, hysteresis, mechanical damage, and traceable verification after a suspected fault. | If a weighing scale still shows 5 kg after the load is removed, manually moving the pointer back to zero does not prove that the scale will measure correctly under load. | The gauge should be removed from measurement duty and returned for assessment. Please provide the service pressure, possible peaks, vibration, temperature, orientation, and installation details so the failure mechanism can be investigated. | A visible symptom should be investigated, not cosmetically corrected. | Draft
Q014 | TCK91320 | Fault diagnosis and claims | Beginner | True/False | If a pressure gauge pointer is not at zero when the system is fully depressurized, adjusting the pointer on site is sufficient to restore measurement accuracy. | TRUE | FALSE |  |  | B | A zero adjustment changes one indication point only. It does not establish that the sensing element, movement, span, linearity, hysteresis, or overload resistance remain correct. | One-point adjustment is not equivalent to calibration across the measuring range. | A clock can be set to the correct time while still running too fast. One correct indication does not prove correct behaviour afterwards. | A zero correction alone is not enough. The cause of the offset and the performance across the range must be checked. | Correcting zero does not prove the measuring function is correct. | Draft
Q015 | TCK91295 | Food-compatible filling fluids | Intermediate | MCQ | A customer specifies one named FDA-compatible fill fluid for a diaphragm-seal system. Another FDA-compatible white oil is available. What is the best engineering response? | Substitute it automatically because both fluids have FDA status | Reject every alternative because only the named trade name can ever work | Compare regulatory status and the fluid properties that affect the complete assembly, then obtain customer acceptance if the named fluid is contractual | Use any oil because the fluid does not contact the process | C | Regulatory suitability is only one requirement. Viscosity, thermal expansion, operating temperature, compatibility, response time, vacuum behaviour, and the customer specification can influence whether an alternative is technically and contractually acceptable. | Hydraulic transmission, viscosity, thermal expansion, material compatibility, and conformity to the specified use. | Two food-grade cooking oils may both be suitable for food contact, but they do not necessarily behave the same way at low temperature or under heat. | Marcol 82 is also an FDA-compatible white oil, but we should confirm that its operating properties and the customer specification are compatible before treating it as an equivalent. | The same approval does not automatically mean identical system behaviour. | Draft
Q016 | TCK91295 | Food-compatible filling fluids | Beginner | True/False | Two fill fluids with the same food-contact approval are automatically interchangeable in every diaphragm-seal assembly. | TRUE | FALSE |  |  | B | Approval addresses a defined regulatory aspect. Interchangeability also depends on physical properties, temperature range, pressure range, vacuum service, material compatibility, and calibration of the filled assembly. | Fluid viscosity and expansion influence force transmission, temperature error, and dynamic response. | Two electrical cables may both be certified, but a thin signal cable cannot automatically replace a power cable simply because both carry an approval mark. | We can propose an alternative approved fluid only after checking the complete application and assembly requirements. | Approval is necessary evidence, not proof of full technical equivalence. | Draft
Q017 | TCK91289 | Pressure-switch setting and hysteresis | Advanced | Advanced Engineering Thinking | A pressure switch is specified with a set point of 20 bar, a setting uncertainty of ±0.45 bar, and hysteresis below 3 bar. Which conclusion is the most responsible? | The switch will always operate at exactly 20.00 bar in both directions | The rising switching point is expected around the specified tolerance, while the reset point also depends on hysteresis and test conditions | The reset point is exactly 17 bar in every installation | Hysteresis and setting uncertainty describe the same error | B | Set-point uncertainty describes the permissible deviation of the switching event under defined conditions. Hysteresis is the difference between switching points when pressure moves in opposite directions. They are different characteristics and both matter when defining an alarm or control band. | Mechanical switching, repeatability, tolerance, hysteresis, direction of approach, and test conditions. | A thermostat may turn heating off near 22°C but not turn it back on until a lower temperature. The difference prevents rapid cycling and is not the same as calibration error. | We should define the required switching direction, allowable trip window, reset requirement, pressure-change rate, and test conditions before confirming that the switch meets the process need. | A set point is not complete without tolerance, direction, and hysteresis. | Draft
Q018 | TCK91289 | Pressure-switch setting and hysteresis | Beginner | True/False | Pressure-switch hysteresis and set-point accuracy describe exactly the same characteristic. | TRUE | FALSE |  |  | B | Set-point accuracy or setting uncertainty concerns how close the switching point is to the target under stated conditions. Hysteresis concerns the difference between switching points on rising and falling pressure. | Tolerance and hysteresis are distinct properties of a switching mechanism. | A door latch can engage at one position and release at another. The gap between those positions is not the same as an error in where the latch was intended to engage. | Please specify both the acceptable trip tolerance and the required reset behaviour. | Do not replace a switching band with a single accuracy number. | Draft
Q019 | TCK91271 | SF6 service and leak risk | Advanced | Advanced Engineering Thinking | A gauge uses apparently compatible stainless-steel wetted parts and the process medium is SF6. Is material compatibility alone enough to approve the application? | Yes, compatible wetted materials settle the complete question | No, leak-tightness, seals, connections, pressure, temperature, regulatory requirements, and the consequence of emissions must also be assessed | Yes, if the gauge has an IP65 enclosure | No, because SF6 cannot be measured mechanically | B | A gauge may be chemically compatible yet unsuitable as a complete installation if the sealing concept, joints, gaskets, permeation, test method, or environmental requirements do not meet the application. Enclosure IP rating concerns ingress into the case, not process-gas leakage out of the pressure boundary. | Chemical compatibility, pressure-boundary integrity, permeation, joint design, and distinction between enclosure protection and process containment. | A waterproof watch may resist rain entering the case, but that says nothing about whether a pressurized gas pipe connection will remain leak-tight. | The wetted materials do not show an obvious limitation, but the customer must validate the complete pressure boundary, sealing method, leakage requirements, operating conditions, and applicable environmental rules. | Material compatibility is only one part of application suitability. | Draft
Q020 | TCK91271 | Ingress protection versus process sealing | Beginner | True/False | An IP65 rating proves that a pressure gauge will not leak process gas through its pressure connection. | TRUE | FALSE |  |  | B | IP ratings describe resistance of the enclosure to ingress of dust and water under defined tests. They do not certify the leak rate of the process connection or internal pressure boundary. | Different barriers protect against different directions and failure modes. | A raincoat can keep water from reaching a person but does not prove that a bottle in the pocket is gas-tight. | IP65 is useful for environmental protection of the case, but process leakage must be assessed separately. | Do not use an enclosure rating as proof of process containment. | Draft
Q021 | TCK91156 | Stainless-steel construction | Intermediate | MCQ | A customer asks for a “fully stainless-steel” pressure gauge. What should sales clarify first? | Only whether the case is shiny | Whether the request concerns the case, movement, pressure element, connection, all wetted parts, or every component | Whether the dial is printed in black | Whether the gauge is filled | B | The phrase “fully stainless” is ambiguous. A standard gauge may combine a 304 stainless case with 316L wetted parts, while another execution may use 316L more broadly. Suitability depends on which components are exposed to the process or environment and why the customer needs the material. | Material grades, wetted versus non-wetted parts, atmospheric corrosion, process corrosion, and functional requirements. | A car described as aluminium may have aluminium body panels but steel axles and fasteners. The statement is meaningless until the relevant components are identified. | Do you require 316L for the process-wetted parts, the external housing, or the complete construction? Please also explain the corrosion or cleaning environment driving the requirement. | Translate broad material language into a component-by-component requirement. | Draft
Q022 | TCK91156 | Stainless-steel construction | Beginner | True/False | If the pressure element and process connection are 316L, the complete pressure gauge is necessarily made entirely from 316L. | TRUE | FALSE |  |  | B | A product can use 316L for wetted parts while the case or other non-wetted components use another stainless grade such as 304. | Material specification must be applied to defined components and exposure conditions. | A stainless-steel kitchen appliance can contain several grades and non-stainless internal parts while still being marketed broadly as stainless steel. | We should provide the material breakdown rather than relying on the phrase “all stainless.” | A product material claim is only meaningful when the covered parts are defined. | Draft
Q023 | TCK91062 | Vacuum measurement with diaphragm seals | Intermediate | MCQ | Can any diaphragm seal measure a -1 to 0 bar vacuum range equally well? | Yes, vacuum is easier than positive pressure | No, diaphragm diameter, stiffness, fill volume, trapped gas, temperature effects, and assembly design influence sensitivity | Yes, if the process connection thread is stainless steel | No diaphragm seal can ever measure vacuum | B | Vacuum measurement provides limited differential force to deflect the diaphragm. A larger and more sensitive diaphragm generally provides better transmission, while a small or stiff diaphragm and a large fill volume can reduce usable sensitivity and increase temperature influence. | Force equals pressure multiplied by effective area; diaphragm stiffness and fill-fluid behaviour influence displacement and transmission. | The same person can move a large lightweight door more easily by pushing far from the hinge than by pushing close to it. Geometry changes the effective response to the applied force. | Vacuum is possible, but we must validate the diaphragm size, construction, fill system, temperature range, and required accuracy. A large connection such as 2 inches may be appropriate where a smaller membrane is not. | Low pressure requires enough effective diaphragm area and a suitable complete assembly. | Draft
Q024 | TCK91062 | Vacuum measurement with diaphragm seals | Beginner | True/False | For a diaphragm seal, increasing the effective diaphragm area can improve sensitivity at low differential pressure. | TRUE | FALSE |  |  | A | For a given pressure, a larger effective area produces a larger total force. The complete design still matters, including diaphragm stiffness, profile, fill fluid, instrument range, and temperature effects. | Force is the product of pressure and effective area. | Snowshoes spread a person’s weight over a larger area. In a diaphragm, area similarly changes the total force generated by a given pressure, although the mechanical objective is different. | For very low pressure or vacuum, we should favour a sufficiently large, sensitive diaphragm and validate the complete filled system. | Pressure alone does not define diaphragm response; effective area matters. | Draft
Q025 | TCK91007 | Sensor and diaphragm-seal system accuracy | Intermediate | MCQ | A low-pressure transmitter is assembled with a diaphragm seal and its declared system accuracy changes from 0.25% to 0.5%. What is the best explanation? | The transmitter electronics necessarily became defective | The added seal and fill system become part of the measurement chain and add influence factors | The number of display digits was reduced | All diaphragm seals always double every error | B | Once a diaphragm seal is assembled, the measured pressure passes through a diaphragm, fill fluid, and mechanical interface before reaching the sensor. Temperature, diaphragm stiffness, fill volume, assembly tolerances, and calibration can add uncertainty or modify response. | Propagation of uncertainty through a complete measurement chain. | Adding an extension cable and converter to a precise sensor can introduce new tolerances. The system specification must include the added elements, not only the original sensor. | The transmitter alone may have a better specification, but the guaranteed performance of the assembled system must include the diaphragm seal and filling effects. | The final specification belongs to the assembled system. | Draft
Q026 | TCK91007 | Sensor and diaphragm-seal system accuracy | Beginner | True/False | The accuracy stated for a bare pressure transmitter always remains unchanged after a diaphragm seal is added. | TRUE | FALSE |  |  | B | The seal assembly introduces additional physical elements and influence quantities. The resulting system must be evaluated and, where necessary, calibrated as an assembly. | Uncertainties and environmental influences accumulate through the measurement chain. | A precise thermometer placed inside a thick protective pocket may respond differently because the pocket becomes part of the thermal path. | Please use the performance declared for the complete assembled system, not only the bare transmitter. | Component accuracy cannot simply be copied to a modified system. | Draft
Q027 | TCK90966 | Supplier competence and welding responsibilities | Advanced | Advanced Engineering Thinking | A customer buys an aluminium part designed to the customer’s dimensions and asks the instrument supplier for welding instructions, but the supplier has no aluminium-welding competence. What is the most responsible response? | Invent generic welding parameters to help the customer quickly | State the competence limitation, provide the controlled drawing, and require the qualified fabricator or design authority to define and qualify the weld procedure | Recommend any available adhesive instead | Approve the customer’s welding method without review | B | Welding quality depends on alloy, thickness, joint design, process, filler, heat input, distortion, service loads, and qualification requirements. Giving unsupported instructions creates false confidence and transfers risk without technical basis. | Competence boundaries, qualified procedures, heat-affected zones, distortion, and responsibility of the design or fabrication authority. | A physician should not prescribe outside their competence simply because the patient needs a fast answer. The correct action is to provide available facts and refer the decision to the qualified specialist. | We can provide the part drawing and interface information, but we do not have the competence to define aluminium welding parameters. The customer or qualified fabricator should apply an approved welding specification. | Knowing when not to improvise is part of engineering judgement. | Draft
Q028 | TCK90953 | Clarification of ambiguous requirements | Intermediate | MCQ | A request mixes a temperature range used on one product family with the code of another family, while the rest of the enquiry refers to different instruments. What should sales do first? | Create a new code using the mixed information | Select the nearest existing product without asking | Clarify the required function, product family, range, and interfaces before quoting | Assume every line refers to the same product | C | A code is the result of a coherent technical definition. Combining attributes from different product families can produce a configuration that has never existed or does not satisfy the intended function. | Requirements engineering, configuration control, functional decomposition, and interface consistency. | Ordering a vehicle by combining an aircraft wing size, a truck engine code, and a car model name does not define a usable product. | The enquiry appears to mix a TBI range with an RT request and other gauge items. Please confirm the required instrument type, function, range, connection, and quantity for each line. | Clarify the function before translating text into a product code. | Draft
Q029 | TCK90953 | Clarification of ambiguous requirements | Beginner | True/False | When a customer request contains contradictory product-family information, choosing the closest code is better than asking for clarification. | TRUE | FALSE |  |  | B | A quick guess may create an apparently complete offer while hiding a fundamental misunderstanding. Clarification prevents incorrect product selection and later rework. | Ambiguous inputs produce unreliable outputs. | A navigation system cannot reliably choose a route when the destination city and postal code point to different places. | Before quoting, we need confirmation of what is being measured and which product family is expected. | Do not make precision-looking selections from contradictory requirements. | Draft
Q030 | TCK90947 | Obsolescence and replacement engineering | Advanced | Advanced Engineering Thinking | An obsolete pressure transmitter has a customer-specific flange, absolute-pressure measurement, food-compatible fill, elevated media temperature, and special calibration. What is the best replacement strategy? | Replace it with any modern transmitter having the same 0 to 5 bar range | Match only the electrical output because mechanics are secondary | Reconstruct the complete functional specification and verify every interface before selecting or engineering a replacement | Tell the customer that obsolete products can never be replaced | C | The measuring range is only one attribute. Process connection geometry, pressure reference, wetted materials, fill fluid, media and ambient temperature, output, connector, accuracy, calibration conditions, labelling, and qualification can all be essential. | Functional equivalence, interface control, thermal effects, pressure reference, and configuration management. | A replacement heart valve cannot be selected only by diameter; material, interface, flow behaviour, and patient conditions must also match. | We can investigate a modern Y91 or PBMN solution, but the customer-specific connection and all operating and certification requirements must first be reconstructed and validated. | Replacement means matching the function and interfaces, not the model name alone. | Draft
Q031 | TCK90947 | Obsolescence and replacement engineering | Beginner | True/False | For an obsolete pressure transmitter, matching the pressure range and output signal is sufficient to guarantee interchangeability. | TRUE | FALSE |  |  | B | Mechanical, thermal, hygienic, metrological, electrical, documentation, and qualification requirements may all prevent direct interchangeability. | A system interface has multiple dimensions, not only range and signal. | A new battery may have the same voltage but still be unusable if its size, connector, chemistry, or current capability differs. | We should compare the complete specification and installation interface before calling a product a replacement. | Two matching headline values do not prove interchangeability. | Draft
Q032 | TCK90921 | Certification scope and evidence | Intermediate | MCQ | A customer asks whether a specific railway gauge has an EN 45545 fire-and-smoke study. No such study was performed for that product. What is the correct response? | Declare compliance because the product is mainly metallic | Use a certificate from another product family | State that the evidence is not available for this product and avoid claiming compliance | Assume compliance because the gauge is already used on trains | C | Compliance is an evidence-based statement tied to a defined product, material set, configuration, test method, and scope. Similar use or predominantly metallic construction does not replace the required product-specific evidence. | Qualification scope, traceability, configuration control, and limits of similarity arguments. | A crash-test result for one vehicle cannot automatically certify another model merely because both vehicles look similar. | We do not have an EN 45545 fire-and-smoke study for this product, so we cannot claim that certification. We can review whether an alternative qualified product exists or whether a new assessment is feasible. | Never convert absence of evidence into a compliance claim. | Draft
Q033 | TCK90921 | Certification scope and evidence | Beginner | True/False | A certificate for one pressure instrument automatically covers another instrument if both are installed in railway applications. | TRUE | FALSE |  |  | B | Certification scope normally depends on the tested or assessed product definition, materials, variants, and intended use. Application similarity alone is insufficient. | Evidence must remain traceable to the configuration being supplied. | A passport belonging to one person cannot prove the identity of another person who travels to the same destination. | We must check whether the exact product or a documented variant is included in the certificate scope. | Certification follows documented scope, not market sector alone. | Draft
Q034 | TCK90920 | Qualification testing and evidence | Advanced | Advanced Engineering Thinking | A nuclear customer asks whether thermal ageing tests exist for a product family. Several tests were performed during the product lifetime, including tests linked to K3 qualification. What should be sent? | Every test ever performed, without checking relevance | Only the newest document, regardless of purpose | The documents whose scope, configuration, ageing profile, and qualification objective match the customer requirement | A general brochure stating that the product is robust | C | A large quantity of test documents is not the same as relevant evidence. The technical response should map the customer requirement to the qualification basis, product configuration, environmental profile, duration, acceptance criteria, and report traceability. | Qualification by representative ageing, scope control, test severity, and evidence mapping. | A medical test result is useful only if it measures the condition being investigated and applies to the patient and time period concerned. | We have several historical tests. The most relevant appear to be those supporting K3 nuclear qualification, but we should first align the customer requirement with the test scope and product configuration. | Select evidence by applicability, not by volume. | Draft
Q035 | TCK90835 | Certification versus unchanged construction | Advanced | MCQ | A legacy product once had FM/CSA approvals. The manufacturer stopped renewing the approvals because of insufficient sales but retained the same construction. Can the current product be sold as FM/CSA approved? | Yes, unchanged construction means the approval remains valid forever | No, approval status depends on current valid certification, surveillance, and scope, not only construction similarity | Yes, if the old certificate is attached | Only if the customer does not ask | B | A product may remain physically similar after certification lapses, but the manufacturer cannot claim a current third-party approval without valid certification and the required ongoing controls. Technical similarity and certification status are different facts. | Third-party conformity assessment, validity periods, surveillance, controlled construction, and legal use of approval marks. | A driver may still have the same skills after a licence expires, but cannot claim to hold a current valid licence. | The product family remains available with similar construction, but FM/CSA approval is no longer available and must not be claimed. | Unchanged design does not preserve an expired certification claim. | Draft
Q036 | TCK90835 | Certification versus unchanged construction | Beginner | True/False | If a manufacturer keeps the same product design, an expired third-party approval can still be presented as current. | TRUE | FALSE |  |  | B | Certification includes current authorization, controlled production, and often surveillance. Design similarity alone is insufficient. | Conformity claims depend on valid evidence and controlled scope. | A calibration certificate from years ago does not prove that an instrument is calibrated today, even if the instrument looks unchanged. | We can describe the construction, but we cannot claim an approval that is no longer maintained. | Technical similarity and certification validity are not interchangeable. | Draft
Q037 | TCK90826 | Chlorinated media and material protection | Advanced | Advanced Engineering Thinking | For a chlorinated medium at about 30°C, a solid Hastelloy C276 lower flange is proposed. A thick PTFE coating may resist the chemical environment better and reduce cost. What must be checked before recommending the coating? | Only the purchase price | Concentration, other chemicals, pressure, temperature, vacuum, coating integrity, permeation, mechanical damage, and process compatibility | Only the flange colour | Nothing, because PTFE is universally compatible | B | Chemical resistance tables are not complete application approvals. A coating can offer excellent resistance but may be limited by temperature, pressure, permeation, adhesion, vacuum, abrasion, cleaning, or damage exposing the substrate. The actual fluid composition matters, including oxidizers and contaminants. | Chemical compatibility, diffusion and permeation, coating adhesion, mechanical integrity, and combined environmental effects. | A rainproof coating protects a wall only while the coating remains continuous. A small cut can expose the material underneath and change the failure mode. | A PTFE coating may be more effective and economical than a solid HC276 flange for this medium, but we need the complete composition and operating conditions before validating it. | Select material systems from the complete environment and failure mode, not from one chemical name. | Draft
Q038 | TCK90826 | Chlorinated media and material protection | Beginner | True/False | Because PTFE has broad chemical resistance, a PTFE-coated flange is automatically suitable for every pressure, temperature, vacuum, and cleaning condition. | TRUE | FALSE |  |  | B | Chemical resistance is only one dimension. Mechanical loading, permeation, coating damage, temperature, pressure, vacuum, and cleaning methods can control suitability. | A barrier works only within its chemical and mechanical limits. | A chemical-resistant glove can still fail if it is punctured, overheated, or used beyond its rated exposure time. | PTFE is promising, but the complete service conditions and coating limitations must be reviewed. | No material should be approved from one attractive property alone. | Draft
Q039 | TCK90797 | Identification from incomplete evidence | Intermediate | MCQ | A replacement request includes partial photographs of an old gauge and only a small visible part of the diaphragm seal. What is the best response? | Present the guessed configuration as an exact identification | Refuse to assist at all | Propose the closest provisional configuration, clearly state the uncertainty, and ask the customer to validate missing interfaces and options | Choose the cheapest standard assembly | C | Photographs may reveal dial size, range, connection orientation, or part of a seal, but hidden options, materials, fill fluid, calibration, and process connection details may remain unknown. The confidence level must be communicated. | Evidence quality, uncertainty management, traceability, and validation of hidden interfaces. | Identifying a machine from one cropped photograph is like identifying a book from a corner of its cover. A hypothesis can be useful, but it is not proof. | Based on the visible features, this configuration appears close, but the diaphragm-seal identification is a guess. Please confirm the process connection, wetted materials, range, filling, options, and nameplate data. | A useful hypothesis must be labelled as a hypothesis. | Draft
Q040 | TCK90797 | Identification from incomplete evidence | Beginner | True/False | If a proposed replacement visually resembles an old instrument, it can be described as an exact match without checking hidden options. | TRUE | FALSE |  |  | B | External appearance does not reveal every material, internal option, fill fluid, calibration, pressure reference, or certification. | Observed similarity is not complete configuration identity. | Two keys can look almost identical while a small hidden cut determines whether either key works in the lock. | We can propose a candidate but need configuration evidence before calling it identical. | Do not turn incomplete evidence into false certainty. | Draft
Q041 | TCK90783 | Standards and connection equivalence | Advanced | Advanced Engineering Thinking | A customer asks for a flanged component to EN 1092, but the proposed supplier’s range does not include EN 1092. What is the correct engineering approach? | Offer the closest-looking flange without comment | Check the requisition and technical requirements, identify the actual dimensional and rating standard needed, and obtain customer approval for any alternative | Assume every DN15 PN50 flange is interchangeable | Change the description in the quotation so the difference is hidden | B | Nominal size and pressure class do not guarantee interchangeability across standards. Facing, drilling, dimensions, pressure-temperature rating, materials, and tube connection details can differ. An alternative requires an explicit technical comparison and acceptance. | Interface compatibility, dimensional standards, pressure-temperature ratings, and configuration control. | A plug and socket may carry the same voltage but remain mechanically incompatible because the pin geometry and standard differ. | The requested EN 1092 execution is not available in this supplier range. We should confirm the true interface requirement and submit the closest alternative with a documented deviation for approval. | Never assume equivalence from similar nominal labels. | Draft
Q042 | TCK90719 | Obsolete railway products and controlled replacement | Intermediate | MCQ | An obsolete railway pressure sensor can still be manufactured for replacement service. Who should decide whether the blocked code is reopened? | Any salesperson who receives the request | The product owner or authorized product-management function after checking configuration and compliance | The customer alone | The warehouse operator | B | A technically manufacturable legacy product can still carry configuration, certification, supply-chain, and liability constraints. Reopening should be a controlled decision by the responsible product authority. | Configuration governance, obsolescence management, traceability, and controlled conformity evidence. | A museum aircraft may still be physically flyable, but returning it to service requires authorization and evidence beyond a mechanic’s ability to assemble it. | The product can technically be manufactured for replacement, but the product manager must authorize reopening the code and confirm the applicable certificates. | Manufacturability does not replace product-governance approval. | Draft
Q043 | TCK90564 | Environmental protection and obsolete variants | Intermediate | MCQ | A customer asks for an old “tropicalized” pressure switch, while the current standard switch has an IP66 enclosure. What should be evaluated? | Assume the old label always means better performance | Compare the original tropicalization purpose with the actual environmental threats and the protection already provided by the current enclosure | Select the old variant because older designs are safer | Ignore humidity because IP ratings cover every environmental effect | B | A tropicalized execution may historically have addressed humidity, condensation, corrosion, insulation, or biological exposure. A modern sealed enclosure can make some measures redundant, but IP66 does not automatically address condensation, internal corrosion, temperature cycling, or every tropical environment. | Ingress protection, humidity, condensation, corrosion, temperature cycling, and changing design solutions over time. | Modern cars no longer need some old external accessories because the protection moved into the base design, but the underlying environmental requirement still needs to be understood. | The former tropicalized variant has been phased out. The standard sealed execution may already cover the need, but please confirm humidity, condensation, temperature cycling, salt, and mounting conditions. | Compare the requirement, not the historical option name. | Draft
Q044 | TCK90564 | Installation orientation of pressure switches | Beginner | True/False | If a pressure switch manual limits installation inclination, mounting the switch horizontally can be assumed acceptable without checking the reference orientation. | TRUE | FALSE |  |  | B | An allowed inclination is defined relative to a specified reference position. The switching mechanism can be influenced by gravity, friction, or mechanical geometry, so orientation must be interpreted from the installation instructions. | Gravity and orientation can shift mechanical switching behaviour. | A spirit level works only when its reference orientation is understood. Rotating it changes what the bubble means. | Please compare the actual mounting to the reference orientation and permitted inclination stated in the instructions. | An angle limit is meaningless without its reference position. | Draft
Q045 | TCK90545 | Incoming inspection and material identification | Advanced | Advanced Engineering Thinking | A package label says a gasket is PTFE, but the received parts visually resemble metal. Two suppliers delivered the same appearance. What is the best next step? | Accept the parts because two suppliers cannot both be wrong | Reject them based only on colour | Quarantine the parts and verify identity against drawings, dimensions, supplier documentation, and suitable material identification methods | Relabel the parts as metal | C | Consistent delivery from multiple suppliers does not prove correctness if both worked from the same incorrect specification or master data. Visual appearance can trigger an investigation but is not normally sufficient material identification. Traceability must connect the article, specification, supplier evidence, and physical item. | Incoming inspection, common-cause error, traceability, material verification, and master-data control. | If two students copy the same wrong source, matching answers do not make the answer correct. Independence of evidence matters. | The appearance conflicts with the description, so the stock should be contained while the article definition, supplier certificates, dimensions, and material are verified. | Repeated consistency can repeat the same error. | Draft
Q046 | TCK90545 | Incoming inspection and material identification | Beginner | True/False | Receiving the same unexpected-looking gasket from two different suppliers proves that the material is correct. | TRUE | FALSE |  |  | B | Both suppliers may have received the same drawing, part description, or master-data error. Independent verification is still required. | Common-cause errors can affect apparently independent sources. | Two thermometers calibrated against the same faulty reference can agree with each other and still both be wrong. | Supplier agreement is useful evidence, but it is not sufficient when the physical item conflicts with the specification. | Agreement is not proof when the sources may share the same error. | Draft
Q047 | TCK90540 | Crush gaskets and reassembly | Intermediate | MCQ | An aluminium gasket in a diaphragm-seal assembly appears permanently flattened after disassembly. What is the most technically correct interpretation? | The gasket is necessarily defective | Crushing may be its intended sealing mechanism, but the opened filled assembly still requires controlled restoration | The gasket can always be reused indefinitely | The shape proves that the process pressure was too high | B | Soft-metal crush gaskets deform plastically to fill surface irregularities and create a seal. Their deformed appearance can therefore be normal after tightening. Reusability and replacement depend on the design, but opening the diaphragm-seal system also releases fill fluid and invalidates the filled measurement chain. | Plastic deformation, contact pressure, sealing, and integrity of a closed hydraulic system. | A metal sealing washer under a fastener changes shape to create contact. Its changed shape can be evidence that it performed its function, not proof that it was manufactured incorrectly. | The gasket deformation can be normal for a crush seal. However, because the filled system was opened, the complete assembly must be refilled and verified or replaced. | Understand the intended failure or deformation mode before judging appearance. | Draft
Q048 | TCK90490 | Certification matrix and scope | Advanced | Advanced Engineering Thinking | A customer requests material, CE, ATEX, PED, PMI, calibration, hydrostatic, leak, and SIL certificates for a mechanical gauge and diaphragm-seal assembly. What should sales do? | Promise the complete list because all certificates concern quality | Reject the entire request because one item is unavailable | Map each requested document to the exact product, function, applicable directive, test scope, and available option, then identify gaps such as SIL | Replace missing certificates with a general ISO 9001 certificate | C | Different certificates answer different questions and may apply to different components. SIL concerns functional safety capability of a safety function and is not automatically relevant to a mechanical gauge. Material certificates, PMI, calibration, pressure tests, and conformity declarations also have distinct scopes and cannot substitute for one another. | Conformity assessment, evidence mapping, component scope, functional safety, and traceability. | A travel folder may require a passport, visa, vaccine record, and ticket. One document cannot replace the others because each proves something different. | We can provide the available certificate options item by item. SIL is not applicable to this mechanical assembly, and each remaining document must be confirmed against the supplied configuration. | Treat certificates as scoped technical evidence, not as a generic paperwork bundle. | Draft
Q049 | TCK90490 | Functional safety certification | Beginner | True/False | A SIL certificate is automatically applicable to every pressure instrument used in a safety-related installation. | TRUE | FALSE |  |  | B | SIL relates to a defined safety function and the capability of the relevant subsystem or device within that function. A mechanical indicator does not become SIL-certified merely because it is installed in a safety environment. | Functional safety depends on architecture, failure behaviour, diagnostics, and defined safety functions. | A fire-exit sign is used in a safety system, but that does not make it a fire alarm detector. Function and evidence must match. | Please define the safety function and the role of the instrument before requesting SIL evidence. | Safety importance does not automatically create SIL applicability. | Draft
Q050 | TCK81801 | Welded pressure connections | Advanced | MCQ | A customer requires a fully welded gauge and diaphragm-seal connection for vibration service. Which configuration principle is most appropriate? | Weld the flange bolts after assembly | Use a connection arrangement designed for welding, such as the specified NPT gauge/seal interface option, and keep bolted flange fasteners as secured mechanical joints | Weld every threaded and bolted component indiscriminately | Use thread tape as a substitute for welding | B | A welded pressure path must be designed around weldable interfaces and materials. Bolts are intended to provide controlled clamping and serviceability; welding them can change material properties, prevent maintenance, and create uncontrolled stresses. | Joint design, load paths, weldability, vibration resistance, and separation of pressure sealing from structural clamping. | A bicycle wheel uses spokes and a hub for different functions. Welding every moving or adjustable connection together does not automatically make the system stronger. | For a welded gauge-to-seal connection, use the intended weldable option. The flange bolts should remain secured fasteners and should not be treated as weld joints. | Strength comes from the correct joint type in the correct location. | Draft
Q051 | TCK81798 | Overpressure diagnosis | Advanced | Advanced Engineering Thinking | Several gauges on a high-pressure gas system show stuck pointers or incorrect readings. The operating range is nominally below the gauge limit, but startup may create pressure spikes above 600 bar. What should be investigated first? | The dial printing colour | The startup sequence, valve operation, pressure-transient measurement, and evidence of mechanical overload | Whether the pointer can be manually pushed back | Whether a larger logo can be printed | B | A fast pressure spike can exceed the elastic capability of the pressure element or disengage movement components even when steady-state pressure is acceptable. Standard process data may miss short transients, so diagnosis may require a suitable fast-response measurement and review of valve events. | Water hammer or gas transients, dynamic overload, inertia, elastic limits, and mechanical gearing. | A road may be safe under steady traffic but damaged by a brief impact from an overloaded vehicle. Average loading does not reveal the peak event. | The failure pattern suggests overload. We should characterize startup and valve-induced transients before selecting a stronger range or protection accessory. | Investigate the event that damages the instrument, not only the normal operating point. | Draft
Q052 | TCK81798 | Overpressure diagnosis | Beginner | True/False | If normal process pressure is below the gauge range, overpressure can be ruled out as the cause of a stuck pointer. | TRUE | FALSE |  |  | B | Short startup, shutdown, blockage, valve, or pulsation events can exceed the normal pressure and may not appear in slow process records. | Peak dynamic conditions can control damage even when average conditions are acceptable. | A computer can overheat during a short workload spike even if its average workload is low. | Please review transients and the full operating cycle before excluding overload. | Normal pressure does not describe every damaging event. | Draft
Q053 | TCK81659 | Traceability and certification | Advanced | Advanced Engineering Thinking | A customer requests ATEX/IECEx evidence, “Traceability Level 2,” and a type-test certificate. What is the best response? | Send every quality certificate available | Clarify what each requested term means in the customer specification, then map it to the conformity declaration, serial traceability, manufacturing records, and available test evidence | Assume ISO 9001 answers every request | State that a serial number is equivalent to all certification | B | Terms such as traceability level can be customer-specific. ATEX/IECEx conformity, serial-number traceability, management-system certification, and type testing are different evidence chains. The response should identify exactly what is required and what can be demonstrated. | Traceability, controlled records, conformity assessment, test evidence, and customer-specific terminology. | A parcel tracking number proves the shipping history, but it does not prove the contents passed a safety test. Different records support different claims. | Please provide the definition of Traceability Level 2 and the required type-test scope. We can then map the request to our declarations, serial records, and available qualification documents. | Clarify undefined evidence terms before promising documents. | Draft
Q054 | TCK71550 | Applicability of electrical standards | Intermediate | MCQ | A customer requests UL certification for a purely mechanical pressure gauge. What is the best first step? | State that every industrial product must have UL | Reject the request without asking anything | Ask which UL standard or hazard they need addressed, because general electrical-safety certification may not apply to a mechanical gauge | Apply an unrelated UL94 material classification to the complete gauge | C | UL is not one universal approval. It includes many standards with defined scopes. A mechanical gauge may fall outside general electrical safety, while a specific material, enclosure, or fire-performance requirement could still be relevant. | Standards applicability, product function, hazard-based assessment, and scope of certification. | Asking whether a bicycle is “ISO certified” is incomplete until the specific standard and requirement are named. | Our gauge is mechanical, so a general electrical UL approval is not normally applicable. Please identify the exact UL standard or performance requirement so we can check it properly. | A standard number and scope are required before compliance can be assessed. | Draft
Q055 | TCK71550 | Applicability of electrical standards | Beginner | True/False | A component material meeting UL94 V-0 automatically makes the complete mechanical instrument UL certified. | TRUE | FALSE |  |  | B | A material flammability classification applies to the tested material and conditions. It does not automatically establish certification of the complete product under an unrelated product standard. | Component evidence has a defined scope and cannot be expanded without justification. | Fire-resistant fabric in a car seat does not certify the entire vehicle for crash safety. | We can state the material classification where applicable, but must not present it as complete product certification. | A component result is not a complete-product approval. | Draft
Q056 | TCK72938 | Field repair and recalibration | Intermediate | MCQ | A customer wants to replace the spring or bellows of a mechanical pressure switch in the field. What is the main technical concern? | The replacement part may have a different colour | The customer may not be able to restore and verify calibration, set point, hysteresis, and safe operation after changing the sensing mechanism | The switch will necessarily become heavier | Field repair is always prohibited for every product | B | The spring or bellows is part of the force balance that determines switching behaviour. Replacing it changes the metrological function and requires controlled adjustment, calibration, and verification. | Force balance, spring rate, effective area, set point, hysteresis, and calibration traceability. | Replacing a piano string changes tension and pitch. Installing the part is not the same as tuning and proving the instrument. | The complete switch can be supplied or serviced through an appropriate facility, but replacing the sensing element in the field without calibration capability is not recommended. | Replacing a functional element requires restoration of the measured function. | Draft